Installation & Operational Qualification (IOQ) Final Report

Project ID: VAL-2023-TSX-001
Equipment: Thermo Scientific TSX Series High-Performance -20°C Manual Defrost Freezer
Model Number: TSX2320FA
Serial Number: TSX-88234-9912
Location: Main Lab - Zone B

1. Installation Qualification (IQ)

Objective: To verify that the TSX2320FA has been delivered and installed according to manufacturer and facility requirements.

| Test ID | Verification Activity | Acceptance Criteria | Actual Result | Status |
| --- | --- | --- | --- | --- |
| IQ-01 | Documentation Review | Manuals, Warranty, and NIST Certificates must be present. | All documents located in the site binder. | PASS |
| IQ-02 | Utility Verification | 115V/60Hz dedicated circuit; NEMA 5-15P plug. | Verified 117.2V at outlet; dedicated breaker. | PASS |
| IQ-03 | Environmental Specs | 4-6" clearance for ventilation; ambient temp <32°C. | 6" rear clearance; Ambient 21.4°C. | PASS |
| IQ-04 | Component Inventory | All shelves, keys, and probe covers included. | 4/4 shelves and 2 keys verified. | PASS |
| IQ-05 | Physical Inspection | Unit must be free of dents, scratches, or seal gaps. | No damage noted during unboxing. | PASS |

2. Operational Qualification (OQ)

Objective: To verify that the freezer subsystems and control interfaces operate within the specified ranges before sample loading.

| Test ID | Verification Activity | Acceptance Criteria | Actual Result | Status |
| --- | --- | --- | --- | --- |
| OQ-01 | Display & Interface | Touch keypad responsive; LED display clear and bright. | Interface fully functional; no dead pixels. | PASS |
| OQ-02 | Door & Seal Integrity | Magnetic gasket must hold "Paper Test" at all points. | Tight seal confirmed around entire perimeter. | PASS |
| OQ-03 | Battery Backup Test | Alarm system stays active for 24hrs on battery. | Battery charged; display stayed on via battery. | PASS |
| OQ-04 | Alarm Triggers | Door Ajar alarm triggers after 3-minute delay. | Alarm activated at 03:00 on the second. | PASS |
| OQ-05 | Temp Verification | Internal display vs. NIST Probe (±1.0°C). | Display: -20.0°C / NIST: -20.2°C. | PASS |
| OQ-06 | Fan/Motor Ops | Internal fans stop when door opens; start when closed. | Verified correct cut-off switch operation. | PASS |

3. Performance Summary (Pre-Load)

Objective: 24-hour stability check at the set point of -20.0°C.

- 24-Hour Mean Temperature: -20.1°C

- Maximum Recorded Temp: -18.9°C (during scheduled defrost cycle/recovery)

- Minimum Recorded Temp: -21.2°C

- Status: PASS (Within ±5.0°C of set point)

4. Deviations & Corrective Actions

- Deviations Noted: None.

- Corrective Actions: N/A.

5. Final Approval

Executive Summary: The Thermo Scientific TSX2320FA (S/N: TSX-88234-9912) has successfully met all predefined acceptance criteria for the IQ and OQ phases. The unit is deemed "Qualified" and is released for Performance Qualification (PQ) / Sample Loading.